## Slide 1: The Art of Flow Cytometry
Sandra Hope, Ph.D.
Microbiology & Molecular Biology Department
College of Life Science
Brigham Young University

## Slide 2: Flow Cytometry
The Flow cytometer
Selecting antibodies
Data interpretation
Control samples
Flow cytometry applications
Variations on a theme
Flow cytometry in the RIC
Notes: Remember – mentored research as well as grad students here. Opportunities to do mentored research – student needs to do some hunting

## Slide 3: Flow Cytometry

## Slide 4: Hydrodynamic Focusing
Sheath Fluid
Sheath Fluid
Sample Core
Flow Cell

## Slide 5: Hydrodynamic Focusing
Flow Cell

## Slide 6: Hydrodynamic Focusing
Laser
readings

## Slide 7: Flow Cytometry
LASER
Forward
Scatter
Detector
(size)
Side
Scatter
Detector
(granularity)

## Slide 8: Flow Cytometry
LASER
Specific Antibody
with
Fluorescent Tag
(protein expression)
Red Fluorescence
Detector

## Slide 9: Flow Cytometry
LASER

## Slide 10: Flow Cytometry
LASER
Forward
Scatter
Detector
Side
Scatter
Detector
Green
Detector
Blue
Detector
Red
Detector
(size)
(granularity)

## Slide 11: Flow Cytometry
LASER
Forward
Scatter
Detector
Side
Scatter
Detector
Green
Detector
Blue
Detector
Red
Detector
Mirrors
(size)
(granularity)

## Slide 12: Fluorochrome Fluorescence
Absorption
Wavelength
Absorption
Wavelength
Emission
Wavelength
Emission
Wavelength
Laser light
Laser light
to PMT
to PMT

## Slide 13: Absorption Wavelengths
FITC
PE
APC

## Slide 14: Absorption Wavelengths
FITC
PE
APC
LASER
LASER

## Slide 15: Emission Wavelengths
FITC
PE
APC

## Slide 16: Emission Wavelengths
FITC
PE
APC

## Slide 17: Flow Cytometry
LASER
Forward
Scatter
Detector
Side
Scatter
Detector
Green
Detector
Blue
Detector
Red
Detector
(size)
(granularity)
Mirrors

## Slide 18: Flow Cytometry
LASER
Forward
Scatter
Detector
Side
Scatter
Detector
Green
Detector
Blue
Detector
Red
Detector
(size)
(granularity)
Mirrors
Filters

## Slide 19: Emission Wavelengths
FITC
PE
APC

## Slide 20: Antibody Selection
Excitation wavelength (laser compatibility)
  488 nm, 633 nm
Emission wavelength (detector compatibility)
  Can the RIC flow cytometer detect the color?
  RIC website, RIC technicians, Dr. Hope
Fluorochrome combinations
  Choose different colors!
  RIC website, RIC technicians, Dr. Hope
Notes:     check with RIC website
    talk to Dr. Hope
    talk to RIC lab technician

## Slide 21: Antibody Selection
http://ricfacility.byu.edu  Equipment  FlowCytometers 
Fluorochromes Detectable by the Cytoflex
Accuri – Student Use
Laser | Fluorescent Channel | Channel Names/Dyes
488 nm - Blue | 533/30 | FL1, FITC, GFP, Alexa-488, SYBR Green, etc.
 | 585/40 | FL2, PE, PI, etc.
 | 670 Long Pass | FL3, PerCP-Cy5.5, PE-Cy7, mCherry, DsRed, etc.
640 nm - Red | 670 Long Pass | FL4, APC, Alexa-647, APC-Cy7, etc.
Laser | Fluorescent Channel | Channel Names/Dyes
488 nm - Blue | 525/40 | BL1, FITC, GFP, Alexa-488, CFSE, Fluo-3, YFP
 | 690/50 | BL2, PerCP, PI, PC5, PC5.5, PerCP-Cy5.5
 | 488/8 | Blue SSC
 | FSC | FSC
638 nm - Red | 780/60 | RL3, APC-H7, APC-A750, APC-Cy7, APC-eFluor780
 | 712/25 | RL2, APC-A700, Alexa-700
 | 660/20 | RL1, APC, Alexa-647, eFluor-660
561 nm - Yellow | 610/20 | YL2, ECD, mCherry, PI, PE-Texas Red, PE-CF594
 | 585/42 | YL1, PE, PI, dTomato, DsRed
 | 690/50 | YL3, PerCP, PI, PC5.5, PerCP-Cy5.5, PC5
 | 780/60 | YL4, PC7, PE-Cy7
405 nm- Violet | 405/10 | Violet SSC (nanoparticle)
 | 450/45 | VL1, Pacific Blue, Brilliant Violet 421, CFP, eFluor-450, V450
 | 525/40 | VL2, CFP-YFP, Krome Orange, AmCyan, V500, Brilliant Violet 510
 | 610/20 | VL3, mCherry, Brilliant Violet 605, Violet 610, Qdot 605
 | Other optional Filters | Optional Colors
 | 780/60 | For Violet laser
 | 660/20 | Brilliant Violet 650, Qdot 655
 | 561/10 | FSC, SSC for yellow laser
Cytoflex – Technician Use
Notes:     check with RIC website
    talk to Dr. Hope
    talk to RIC lab technician

## Slide 22: Flow Cytometry Data
FSC
SSC

## Slide 23: Flow Cytometry Data
GFP
PE
GFP

## Slide 24: Flow Cytometry Data
GFP
PE
GFP
Scatterplots
&
Histograms

## Slide 25: Flow Cytometry Data
GFP
PE
GFP
Regions

## Slide 26: Flow Cytometry Data
FSC
GFP
SSC
Color Gates

## Slide 27: Flow Cytometry Data
FSC
GFP
SSC
Applying
a Gate to
another graph

## Slide 28: Flow Cytometry Data
Data Interpretation
- Identification of
populations

## Slide 29: Flow Cytometry Data
Identification of
Populations
using color gates

## Slide 30: Multicolor Flow Cytometry
Identification of Populations
using applied gates

## Slide 31: Multicolor Flow Cytometry
Immunophenotyping Utilizing 6 Color Flow Cytometry
2012,Ward Medical Laboratory, Volume 22, Number 3
http://www.wardelab.com/22-3.html
T cell Analysis – 6 colors

## Slide 32: Fluorescence Intensity Settings
LASER
Filter for
Green Fluorescence
10
1
10
2
10
3
10
4
55
111
167
223
Counts
R1
10
0
0
Green Fluorescence
Negative Control Sample
with proper voltage settings on
the PMT detecting green
PMT
General
light
scatter
General light scatter
peak should be here
Negative Control Sample
What does the technician do with your control samples?
Cytoflex Cytometer

## Slide 33: Fluorescence Intensity Settings
LASER
Filter for
Green Fluorescence
Voltage
controls
sensitivity
10
1
10
2
10
3
10
4
55
111
167
223
Counts
R1
10
0
0
Green Fluorescence
Negative Control Sample
with voltage set too high
on PMT for green
Sample should
peak here
PMT
General
light
scatter
Negative Control Sample
Cytoflex Cytometer

## Slide 34: Fluorescence Intensity Settings
LASER
Filter for
Green Fluorescence
Voltage
controls
sensitivity
10
1
10
2
10
3
10
4
55
111
167
223
Counts
R1
10
0
0
Green Fluorescence
Negative Control Sample
with voltage set extremely high
on PMT for green
Sample should
peak here
PMT
General
light
scatter
Negative Control Sample
Cytoflex Cytometer

## Slide 35: Fluorescence Intensity Settings
LASER
Filter for
Green Fluorescence
10
1
10
2
10
3
10
4
55
111
167
223
Counts
R1
10
0
0
Red Fluorescence
Positive Control Sample
with proper voltage settings on
the PMT detecting red
PMT
Any unstained cells
In the sample are
measured here
Any cells stained with green-labelled antibody
are measured here
Single-Color Control Sample:
(check red sample for a red reading)
Cytoflex Cytometer

## Slide 36: Fluorescence Overlap
LASER
Filter for
Green Fluorescence
10
1
10
2
10
3
10
4
55
111
167
223
Counts
R1
10
0
0
Green Fluorescence
Red Control Sample (reading in green should be negative)
Some red fluorescence is able to
cross through the green filter
PMT
A small amount of red can pass through the green filter
Single-Color Control Sample:
(check red sample for a green reading)
Cytoflex Cytometer

## Slide 37: Fluorescence Overlap
Spectral Overlap
FITC
PE
APC

## Slide 38: Compensation Matrix
For proper compensation matrix calculations,
each experiment needs:
  Negative control
  Single color positive controls = 1 for each color

## Slide 39: Fluorescence Overlap
  Requires Compensation (digital threshold)
LASER
Filter for
Green Fluorescence
10
1
10
2
10
3
10
4
55
111
167
223
Counts
R1
10
0
0
Green Fluorescence
Red Control Sample (reading in green should be negative)
Some red fluorescence is able to
cross through the green filter
PMT
A small amount of red can pass through the green filter

## Slide 40: Compensation subtracts false readings
LASER
Filter for
Green Fluorescence
10
1
10
2
10
3
10
4
55
111
167
223
Counts
R1
10
0
0
Green Fluorescence
Red Control Sample (reading in green should be negative)
Instrument controls compensate for
“fluorescence overlap” of red
PMT
A small amount of red can pass through the green filter

## Slide 41: Flow Cytometry Applications
Propidium Iodide stains DNA

## Slide 42: Flow Cytometry Applications
Cell Cycle
Division
Apoptosis
Nuclear DNA content

## Slide 43: Flow Cytometry Applications
Cell Death
Live = esterase activity)
Dead = DNA binding
(compromised membranes)
Apoptosis = Annexin V
Dead = PI

## Slide 44: Flow Cytometry Applications
Detecting Proteins
in Supernatants
pStat5
pStat3
Beads
Cytokine Binding
Negative control
Positive control
Cytokine panels
Phosphorylation

## Slide 45: Variations on a Theme
Amnis ImageStream
Imaging Cytometer
Dark-field Microscopy
Camera

## Slide 46: FACS Fluorescence – Activated Cell Sorting
Laser
to
PMT’s
Sort 1
Sort 2
Waste
Charge droplet
as it leaves
deflection
plate
deflection
plate
+
-
Variations on a Theme

## Slide 47: FACS Data
A
B
C
D
E
G
F

## Slide 48: RIC Flow Cytometry Equipment
Cytoflex Flow Cytometer
4 laser, 13 color
Technician-run
Fall Semester Flow Class
Accuri Flow Cytometer
2 laser, 4 color
Fixed-voltage
Student-run

## Slide 49: RIC Facility website http://ricfacility.byu.edu

## Slide 50: RIC Facility Scheduler

## Slide 51: RIC Facility Scheduler http://ricfacility.byu.edu

## Slide 52: The Art of Flow Cytometry